{"text_format":"2","workbook":"xls","n_sheets":26,"sheets":["Data","Language","Version","Menu","QAL1","QAL1 S-AMS","QAL1 Selectivity","QAL1 other sources","QAL1 NumDif","QAL1 S-AMS (zero)","QAL2","QAL2 calibration","QAL2 functionality","VCR-test","QAL3","CUSUM NOx","CUSUM NO","CUSUM NO2","CUSUM O2","CUSUM H2O","AST","AST parallel measurements 1","AST functionality 1","AST parallel measurements 2","AST functionality 2","Blad1"],"not_rendered":{"charts":8,"images":19,"macros":1,"external_links":1,"defined_names":3,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"run_cost":"one","rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900","formulas":"cached_only"}

{"sheet":"Data","index":0,"n_rows":73,"n_cols":17,"rows_kept":73,"cut":false,"hidden":true,"merged":[],"flags":["formulas_cached_only"]}
["Optie niet beschikbaar; verwijder eerst beveiliging van dit werkblad","2.1","Te gebruiken tot","2015-01-01",true]
[null,"NOx-meting","O2-meting","Vochtmeting","Rookgastemperatuurmeting","Rookgasdrukmeting","Monsternamemethode","NOx","NO","NO2","O2","H2O","Jaar","Middelingstijd","NOx","NO","NO2"]
["keuze",2,2,2,1,1,1,1,1,1,1,1,2,1,2,1,1]
["selectie","Extractieve NO- + NO2-meting in droog rookgas","In-situ O2-meting","Berekening op T-scrubber","Thermokopppel","Druksensor","Puntmeting","ppm","ppm","ppm","vol%","vol%","Jaar 2: 2017","0,5 uur","","vppm","vppm"]
[null,"Extractieve NO-meting in droog rookgas","Extractieve O2-meting in droog rookgas","Extractieve IR-meting","Thermokopppel","Druksensor","Puntmeting","ppm","ppm","ppm","vol%","vol%","Jaar 1: 2016","0,5 uur","","vppm","vppm"]
[null,"Extractieve NO- + NO2-meting in droog rookgas","In-situ O2-meting","Berekening op T-scrubber","PT100","Constante","Lijnmeting","mg/Nm3","mg/Nm3","mg/Nm3",null,null,"Jaar 2: 2017","1 uur","","mg/Nm3","mg/Nm3"]
[null,"Extractieve NOx-meting in droog rookgas","In-situ O2-meting (T gecorrigeerd)","Berekend op basis van verbrandingsparameters",null,null,"Netwerkmeting",{"$nulls":7},2,1,1]
[null,"In-situ NO-meting","In-situ O2-meting (T en p gecorrigeerd)","Constante",null,null,"Traverse meting",{"$nulls":7},"","",""]
[null,"In-situ NO-meting (T gecorrigeerd)","Extractieve O2-meting in nat rookgas",{"$nulls":11},"","",""]
[null,"In-situ NO-meting (T en p gecorrigeerd)",{"$nulls":12},"","",""]
[null,"In-situ NO- + NO2-meting"]
[null,"In-situ NO- + NO2-meting (T gecorrigeerd)"]
[null,"In-situ NO- + NO2-meting (T en p gecorrigeerd)"]
[null,"Extractieve NO-meting in nat rookgas"]
[null,"Extractieve NO- + NO2-meting in nat rookgas"]
[null,"Extractieve NOx-meting in nat rookgas"]
[]
["NOx-concentratie",0]
["NOx-meting",false]
["NO-meting",true]
["NO2-meting",true]
["Vochtmeting",true]
["Extractief",true]
["In situ",true]
[]
["Vochtcorrectie",false,true]
["Drukcorrectie",false,true,true]
["Temperatuurcorrectie",false,true,true]
[]
["Eigen data voor S AMS",false]
["Eigen data voor S AMS zero",false]
["Eigen selectiviteitsdata",false]
["Eigen gegevens voor overige onzekerheidsbronnen",false]
[]
["Programma ontwikkeling",false]
["QAL1",false,"QAL1"]
["S-AMS ",false,"QAL1 S-AMS"]
["Selectiviteit",false,"QAL1 Selectivity"]
["Overige onzekerheidsbronnen",false,"QAL1 other sources"]
["Numerieke differentiatie",false,"QAL1 NumDif"]
["S-AMS nul ",false,"QAL1 S-AMS (zero)"]
["QAL2",false,"QAL2"]
["Kalibratie",false,"QAL2 calibration"]
["Functionele testen",false,"QAL2 functionality"]
["Wekelijkse controle geldig kalibratiebereik (VCR-test)",false,"VCR-test"]
["QAL3",false,"QAL3"]
["Cusum NO-meting",false,"Cusum NO","Cusum NOx"]
["Cusum NO2-meting",false,"Cusum NO2"]
["Cusum O2-meting",false,"Cusum O2"]
["",false,"Cusum H2O"]
["AST",false,"AST"]
["Parallelmetingen",false,"AST parallel measurements 2","AST parallel measurements 1"]
["Functionele testen",false,"AST functionality 2","AST functionality 1"]
["Versiebeheer",false,"Version"]
["Taal",false,"Language"]
["Data",false,"Data"]
[]
[]
[]
[]
[]
[]
[null,"Testwaarde","Volle schaal"]
[]
["Nox","","","",""]
["NO",0,0,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
["NO2",0,0,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
["O2","","","",""]
["H2O","","","",""]
[]
[]
[]
[{"$nulls":4},{"$error":"#DIV/0!"}]

{"sheet":"Language","index":1,"n_rows":284,"n_cols":3,"rows_kept":284,"cut":false,"hidden":true,"merged":[],"flags":["formulas_cached_only"]}
[1,"Nederlands","Engels"]
["Taal","Taal","Language"]
["Nederlands","Nederlands","Dutch"]
["Engels","Engels","English"]
[""]
[""]
[""]
["Optie niet beschikbaar; verwijder eerst beveiliging van dit werkblad","Optie niet beschikbaar; verwijder eerst beveiliging van dit werkblad","Option not available; first unprotect this sheet"]
["Versie","Versie","Version"]
["Versiebeheer","Versiebeheer","Version management"]
["Veranderingen t.o.v. vorige versie","Veranderingen t.o.v. vorige versie","Changes since last version"]
["Te gebruiken tot","Te gebruiken tot","Valid until"]
["Programma ontwikkeling","Programma ontwikkeling","Program development"]
["Ondanks de zorgvuldigheid waarmee dit werkblad is opgesteld, zijn fouten niet uit te sluiten.","Ondanks de zorgvuldigheid waarmee dit werkblad is opgesteld, zijn fouten niet uit te sluiten.","Although this spreadsheet has been carefully developed, mistakes can not be excluded."]
["NEN-EN 14181 voor NOx-meetsystemen","NEN-EN 14181 voor NOx-meetsystemen","EN14181 for NOx measurement systems"]
["Resultaten QAL1 voor NOx-meetsystemen","Resultaten QAL1 voor NOx-meetsystemen","Results QAL1 for NOx measurement systems"]
["Berekening S-AMS van het NOx-meetsysteem","Berekening S-AMS van het NOx-meetsysteem","Calculation of S-AMS of the NOx measurement system"]
["Berekening S-AMS (zero) van het NOx-meetsysteem","Berekening S-AMS (zero) van het NOx-meetsysteem","Calculation of S-AMS (zero) of the NOx measurement system"]
["Berekening onzekerheid van het NOx-meetsysteem t.g.v. overig bronnen ","Berekening onzekerheid van het NOx-meetsysteem t.g.v. overig bronnen ","Calculation uncertainty of the NOx measurement system due to other uncertainty sources "]
["Berekening selectiviteit van het NOx-meetsysteem","Berekening selectiviteit van het NOx-meetsysteem","Calculation of the selectivity of the NOx measurement system"]
["Berekening van de totale onzekerheid m.b.v. numerieke differentiatie","Berekening van de totale onzekerheid m.b.v. numerieke differentiatie","Calculation of the total uncertainty by numeric differentiation"]
["Kalibratie van het NOx-meetsysteem","Kalibratie van het NOx-meetsysteem","Calibration of the NOx measurement system"]
["AST-procedure van het NOx-meetsysteem ","AST-procedure van het NOx-meetsysteem ","Annual Surveillance Test of the NOx measurement system"]
["Vereenvoudigde methode voor de QAL1 voor (bestaande) NOx-meetsystemen","Vereenvoudigde methode voor de QAL1 voor (bestaande) NOx-meetsystemen","Simplified QAL1 procedure for (existing) NOx measurement systems"]
["Resultaten QAL2 voor NOx-meetsystemen","Resultaten QAL2 voor NOx-meetsystemen","Results QAL2 for NOx measurement systems"]
["QAL2 - functionele testen","QAL2 - functionele testen","QAL2 - functionality tests"]
["AST - functionele testen","AST - functionele testen","AST - functionality tests"]
["Resultaten QAL3-procedure van het NOx-meetsysteem","Resultaten QAL3-procedure van het NOx-meetsysteem","Results QAL3 procedure of the NOx measurement system "]
["Wekelijkse controle geldig kalibratiebereik","Wekelijkse controle geldig kalibratiebereik","Weekly test on the valid calibration range"]
["Resultaten AST voor NOx-meetsystemen","Resultaten AST voor NOx-meetsystemen","Results AST for NOx measurement systems"]
["Cusum voor NOx-analyser","Cusum voor NOx-analyser","Cusum for NOx analyser"]
["Cusum voor NO-analyser","Cusum voor NO-analyser","Cusum for NO analyser"]
["Cusum voor NO2-analyser","Cusum voor NO2-analyser","Cusum for NO2 analyser"]
["Cusum voor O2-analyser","Cusum voor O2-analyser","Cusum for O2 analyser"]
["Cusum voor H2O-analyser","Cusum voor H2O-analyser","Cusum for H2O analyser"]
["Specificatie meetsysteem","Specificatie meetsysteem","Specification of the measurement system"]
["NOx-meetsysteem (AMS)","NOx-meetsysteem (AMS)","NOx measurement system (AMS)"]
["NOx-meting","NOx-meting","NOx-measurement"]
["NO-meting","NO-meting","NO-measurement"]
["NO2-meting","NO2-meting","NO2-measurement"]
["O2-meting","O2-meting","O2-measurement"]
["Vochtmeting","Vochtmeting","H2O-measurement"]
["Rookgastemperatuurmeting","Rookgastemperatuurmeting","Flue gas temperature measurement"]
["Rookgasdrukmeting","Rookgasdrukmeting","Flue gas pressure measurement"]
["Extractief","Extractief","Extractive"]
["In situ","In situ","In situ"]
["Extractieve NO-meting in droog rookgas","Extractieve NO-meting in droog rookgas","Extractive NO measurement in dry flue gas"]
["Extractieve NO- + NO2-meting in droog rookgas","Extractieve NO- + NO2-meting in droog rookgas","Extractive NO and NO2 measurement in dry flue gas"]
["Extractieve NOx-meting in droog rookgas","Extractieve NOx-meting in droog rookgas","Extractive NOx measurement in dry flue gas"]
["In-situ NO-meting","In-situ NO-meting","Insitu NO measurement"]
["In-situ NO-meting (T gecorrigeerd)","In-situ NO-meting (T gecorrigeerd)","Insitu NO measurement (T corrected)"]
["In-situ NO-meting (T en p gecorrigeerd)","In-situ NO-meting (T en p gecorrigeerd)","Insitu NO measurement (T and p corrected)"]
["In-situ NO- + NO2-meting","In-situ NO- + NO2-meting","Insitu NO + NO2 measurement"]
["In-situ NO- + NO2-meting (T gecorrigeerd)","In-situ NO- + NO2-meting (T gecorrigeerd)","Insitu NO + NO2 measurement (T corrected)"]
["In-situ NO- + NO2-meting (T en p gecorrigeerd)","In-situ NO- + NO2-meting (T en p gecorrigeerd)","Insitu NO + NO2 measurement (T and p corrected)"]
["Extractieve NO-meting in nat rookgas","Extractieve NO-meting in nat rookgas","Extractive NO measurement in wet flue gas"]
["Extractieve NO- + NO2-meting in nat rookgas","Extractieve NO- + NO2-meting in nat rookgas","Extractive NO and NO2 measurement in wet flue gas"]
["Extractieve NOx-meting in nat rookgas","Extractieve NOx-meting in nat rookgas","Extractive NOx measurement in wet flue gas"]
["Extractieve O2-meting in droog rookgas","Extractieve O2-meting in droog rookgas","Extractive O2 measurement in dry flue gas"]
["In-situ O2-meting","In-situ O2-meting","Insitu O2 measurement"]
["In-situ O2-meting (T gecorrigeerd)","In-situ O2-meting (T gecorrigeerd)","Insitu O2 measurement (T corrected)"]
["In-situ O2-meting (T en p gecorrigeerd)","In-situ O2-meting (T en p gecorrigeerd)","Insitu O2 measurement (T and p corrected)"]
["Extractieve O2-meting in nat rookgas","Extractieve O2-meting in nat rookgas","Extractive O2 measurement in wet flue gas"]
["Extractieve IR-meting","Extractieve IR-meting","Extractive IR measurement"]
["Berekening op T-scrubber","Berekening op T-scrubber","Calculation based on scrubber temperature"]
["Berekend op basis van verbrandingsparameters","Berekend op basis van verbrandingsparameters","Calculation based on combustion parameters"]
["Constante","Constante","Constant"]
["N.v.t.","N.v.t.","Not relevant"]
["Thermokopppel","Thermokopppel","Thermocouple"]
["PT100","PT100","PT100"]
["Druksensor","Druksensor","Sensor"]
["Meting","Meting","Measurement"]
["Berekening onzekerheid meetinstrument","Berekening onzekerheid meetinstrument","Calculation uncertainty analyser"]
["Invoer noodzakelijk","Invoer noodzakelijk","Input required"]
["Ja","Ja","Yes"]
["Nee","Nee","No"]
["Rookgas-temperatuur","Rookgas-temperatuur","Flue gas temperature"]
["Absolute rookgas druk","Absolute rookgas druk","Flue gas pressure (absolute)"]
["Testwaarde","Testwaarde","Test value"]
["Eenheid","Eenheid","Unit"]
["Volle schaal","Volle schaal","Full scale"]
["Concentratie QAL3-medium","Concentratie QAL3-medium","Concentration QAL3 medium"]
["Onzekerheid QAL3-medium (%) *","Onzekerheid QAL3-medium (%) *","Uncertainty QAL3 medium (%) *"]
["Concentratie eenheid QAL3-medium","Concentratie eenheid QAL3-medium","Unit concentration QAL3 medium"]
["Data QAL3 procedure","Data QAL3 procedure","Data QAL3 procedure"]
["S-AMS (%) *","S-AMS (%) *","S-AMS (%) *"]
["S-AMS nul (%) *","S-AMS nul (%) *","S-AMS zero (%) *"]
["S-AMS *","S-AMS *","S-AMS *"]
["S-AMS nul *","S-AMS nul *","S-AMS zero *"]
["Onzekerheid t.g.v. overige bronnen *","Onzekerheid t.g.v. overige bronnen *","Uncertainty due to other sources *"]
["Onzekerheid t.g.v. overige bronnen (%) *","Onzekerheid t.g.v. overige bronnen (%) *","Uncertainty due to other sources (%) *"]
["Onzekerheid t.g.v. selectiviteit (%) *","Onzekerheid t.g.v. selectiviteit (%) *","Uncertainty due to selectivity (%) *"]
["Onzekerheid t.g.v. converterrendement (%) *","Onzekerheid t.g.v. converterrendement (%) *","Uncertainty due to converter efficiency (%) *"]
["Onzekerheid in aanname NO2-aandeel (%) *","Onzekerheid in aanname NO2-aandeel (%) *","Uncertainty due to assumption of NO2 contribution (%) *"]
["Totale onzekerheid meting (%) *","Totale onzekerheid meting (%) *","Total uncertainty of measurement (%) *"]
["OPMERKINGEN","OPMERKINGEN","REMARKS"]
["Voer geschatte waarden in voor:","Voer geschatte waarden in voor:","Enter estimated values for:"]
["Geen","Geen","None"]
["Voer volle schaal in voor:","Voer volle schaal in voor:","Enter full scale for:"]
["Voer eenheid in voor:","Voer eenheid in voor:","Enter unit for:"]
["Voer onzekerheid QAL3-medium in voor:","Voer onzekerheid QAL3-medium in voor:","Enter uncertainty QAL3 medium for:"]
["Foutieve zuurstofconcentratie: >21 vol% droog rookgas.","Foutieve zuurstofconcentratie: >21 vol% droog rookgas.","Wrong oxygen concentration: >21 vol% dry flue gas."]
["Voer data in voor:","Voer data in voor:","Enter data for:"]
["Voer de emissiegrenswaarde in","Voer de emissiegrenswaarde in","Enter the emission limit value"]
["Voer de contentratie profiel gegevens in","Voer de contentratie profiel gegevens in","Enter the concentration profile data"]
["NOx-concentratie","NOx-concentratie","NOx-concentration"]
["Rookgascondities","Rookgascondities","Flue gas parameters"]
["Geschatte O2-concentratie","Geschatte O2-concentratie","Estimated O2-concentration"]
["Geschat aandeel NO2","Geschat aandeel NO2","Estimated percentage NO2"]
["Geschatte H20-concentratie","Geschatte H20-concentratie","Estimated H2O-concentration"]
["Geschatte rookgastemperatuur","Geschatte rookgastemperatuur","Estimated flue gas temperature"]
["Geschatte absolute druk van het rookgas","Geschatte absolute druk van het rookgas","Estimated absolute flue gas pressure"]
["Referentie zuurstofconcentratie","Referentie zuurstofconcentratie","Reference oxygen concentration"]
["NOx-concentratie","NOx-concentratie","NOx concentration"]
["O2-concentratie","O2-concentratie","O2 concentration"]
["NOx-concentratie bij standaard condities","NOx-concentratie bij standaard condities","NOx concentration at standard conditions"]
["Emissiegrenswaarde","Emissiegrenswaarde","Emisson limit value"]
["Testwaarde bij standaard condities","Testwaarde bij standaard condities","Test value at standard conditions"]
["Resultaten concentratieprofielmeting","Resultaten concentratieprofielmeting","Results concentration profile measurement"]
["Ontbrekende gegevens concentratieprofiel","Ontbrekende gegevens concentratieprofiel","Missing concentration profile data"]
["Hoogste herleide NOx-concentratie","Hoogste herleide NOx-concentratie","Highest observed NOx concentration at standard conditions"]
["Laagste herleide NOx-concentratie","Laagste herleide NOx-concentratie","Lowest observed NOx concentration at standard conditions"]
["Aantal meetpunten AMS","Aantal meetpunten AMS","Number of AMS sampling points"]
["Aantal meetpunten SRM","Aantal meetpunten SRM","Number of SRM sampling points"]
["Geschatte onzekerheid t.g.v. niet representatieve monstername AMS *","Geschatte onzekerheid t.g.v. niet representatieve monstername AMS *","Estimated uncertainty due to non-representative sampling of AMS"]
["Betrouwbaarheidsinterval","Betrouwbaarheidsinterval","Confidence interval"]
["Meetonzekerheid herleide NOx-concentratie *","Meetonzekerheid herleide NOx-concentratie *","Uncertainty NOx concentration at standard conditions"]
["Berekend betrouwbaarheidsinterval","Berekend betrouwbaarheidsinterval","Calculated confidence interval"]
["Vereist betrouwbaarheidsinterval","Vereist betrouwbaarheidsinterval","Required confidence interval"]
["Resultaat toetsing betrouwbaarheidsinterval","Resultaat toetsing betrouwbaarheidsinterval","Result compliance test confidence interval"]
[" *: als standaard deviatie"," *: als standaard deviatie","*: as standard deviation"]
["vol% droog rookgas","vol% droog rookgas","vol% dry flue gas"]
["vol% nat rookgas","vol% nat rookgas","vol% wet flue gas"]
[" mg/Nm3 (als NO2) droog rookgas"," mg/Nm3 (als NO2) droog rookgas","mg/Nm3 (as NO2) dry flue gas"]
["mg/Nm3 (als NO2) droog rookgas bij","mg/Nm3 (als NO2) droog rookgas bij","mg/Nm3 (as NO2) dry flue gas at"]
["als NO2","als NO2","as NO2"]
["als NO","als NO","as NO"]
["vol% O2","vol% O2","vol% O2"]
["Voldoet niet","Voldoet niet","Non-compliant"]
["Voldoet","Voldoet","Compliant"]
["keuze","keuze","choice"]
["selectie","selectie","Selected item"]
["Vochtcorrectie","Vochtcorrectie","Moisture corrrection"]
["Drukcorrectie","Drukcorrectie","Pressure correction"]
["Temperatuurcorrectie","Temperatuurcorrectie","Temperature correction"]
["Eigen data voor S AMS","Eigen data voor S AMS","User specific S AMS data"]
["Eigen data voor S AMS zero","Eigen data voor S AMS zero","User specific S AMS data zero"]
["Eigen selectiviteitsdata","Eigen selectiviteitsdata","User specific selectivity data"]
["Eigen gegevens voor overige onzekerheidsbronnen","Eigen gegevens voor overige onzekerheidsbronnen","User specific data for other uncertainties"]
["Default waarden","Default waarden","Default values"]
["Waarden als % van de volle schaal","Waarden als % van de volle schaal","Values expressed as % of the full scale"]
["Instabiliteit (drift)","Instabiliteit (drift)","Instability (drift)"]
["Gevoeligheid voor omgevingsdruk","Gevoeligheid voor omgevingsdruk","Sensitivity for ambient pressure"]
["Gevoeligheid voor omgevingstemperatuur","Gevoeligheid voor omgevingstemperatuur","Sensitivity for ambient temperature"]
["Gevoeligheid voor variaties netspanning","Gevoeligheid voor variaties netspanning","Sensitivity for voltage"]
["Instrument specifieke waarden","Instrument specifieke waarden","Instrument specific values"]
["Variaties","Variaties","Variations"]
["Onzekerheid *","Onzekerheid *","Uncertainty *"]
["Berekening S-AMS","Berekening S-AMS","Calculation of S-AMS"]
["% TW","% TW","% TV"]
["% VS","% VS","% FS"]
["% testwaarde","% testwaarde","% test value"]
["% volle schaal","% volle schaal","% full scale"]
["Berekening S-AMS (nul)","Berekening S-AMS (nul)","Calculation of S-AMS (zero)"]
["Afwijking lineariteit *","Afwijking lineariteit *","Linearity *"]
["Herhaalbaarheid *","Herhaalbaarheid *","Repeatability *"]
["Verliezen in monsternamesysteem *","Verliezen in monsternamesysteem *","Losses in sampling system *"]
["Overige onzekerheidsbronnen","Overige onzekerheidsbronnen","Other uncertainty sources"]
["Variaties rookgassamenstelling","Variaties rookgassamenstelling","Variations flue gas composition"]
["Component","Component","Substance"]
["Default selectiviteit","Default selectiviteit","Default selectivity"]
["Min","Min","Min"]
["Max","Max","Max"]
["Gemiddeld","Gemiddeld","Average"]
["Instrument specifieke selectiviteit","Instrument specifieke selectiviteit","Instrument specific selectivity"]
["Selectiviteit","Selectiviteit","Selectivity"]
["Spreadsheet methode","Spreadsheet methode","Spreadsheet method"]
["Parameter","Parameter","Variable"]
["Vocht NOx","Vocht NOx","Moisture NOx"]
["Vocht O2","Vocht O2","Moisture O2"]
["Representativiteit","Representativiteit","Representativity"]
["Waarde","Waarde","Value"]
["St. dev.","St. dev.","St.dev."]
["Relatief?","Relatief?","Relative?"]
["Kalibratie","Kalibratie","Calibration"]
["Kalibratielijn gebruikt tijdens kalibratie","Kalibratielijn gebruikt tijdens kalibratie","Calibration curve used during calibration"]
["Kalibratielijn gebruikt tijdens AST","Kalibratielijn gebruikt tijdens AST","Calibration curve used during AST"]
["Helling","Helling","Slope"]
["Intercept","Intercept","Intercept"]
["Standaard referentie methode (SRM)","Standaard referentie methode (SRM)","Standard Reference Method (SRM)"]
["Meetinstantie","Meetinstantie","Test laboratory"]
["Meetplaats","Meetplaats","Measurement point"]
["Monsternamemethode","Monsternamemethode","Sampling procedure"]
["Puntmeting","Puntmeting","Point sampling"]
["Lijnmeting","Lijnmeting","Line sampling"]
["Netwerkmeting","Netwerkmeting","Network sampling"]
["Traverse meting","Traverse meting","Traverse sampling"]
["Aantal meetpunten","Aantal meetpunten","Number of sampling points"]
["NOx-meetmethode","NOx-meetmethode","NOx-measurement method"]
["O2-meetmethode","O2-meetmethode","O2-measurement method"]
["Kalibratiepunten","Kalibratiepunten","Calibration points"]
["AST meetpunten","AST meetpunten","AST points"]
["Nummer","Nummer","Number"]
["Datum","Datum","Date"]
["Starttijd","Starttijd","Start"]
["Eindtijd","Eindtijd","End"]
["Meetwaarde","Meetwaarde","Value"]
["Verschil","Verschil","Difference"]
["Regressie-lijn","Regressie-lijn","Regression line"]
["Aantal parallelmetingen","Aantal parallelmetingen","Number of parallel measurements"]
["Parallelmetingen","Parallelmetingen","Parallel measurements"]
["AMS","AMS","AMS"]
["SRM","SRM","SRM"]
["Geldig kalibratiebereik","Geldig kalibratiebereik","Valid calibration range"]
["Nieuwe kalibratie lijn","Nieuwe kalibratie lijn","New calibration curve"]
["Kalibratie lijn","Kalibratie lijn","Calibration curve"]
["Variabiliteitstest","Variabiliteitstest","Test for variability"]
["Variabiliteit","Variabiliteit","Variability"]
["Te weinig metingen!","Te weinig metingen!","Number of measurements too small!"]
["Metingen niet verdeeld over 3 dagen in een periode van 4 weken!","Metingen niet verdeeld over 3 dagen in een periode van 4 weken!","Measurements not performed on three days within four weeks!"]
["Dag","Dag","Day"]
["AST resultaten","AST resultaten","AST Results"]
["Test voor validiteit kalibratie lijn","Test voor validiteit kalibratie lijn","Test for validity of calibration curve"]
["Valide","Valide","Valid"]
["Niet-valide","Niet-valide","Invalid"]
["Numerieke differentiatie","Numerieke differentiatie","Numeric differentation"]
["Functionele testen","Functionele testen","Functionality tests"]
["Nul","Nul","Zero"]
["Span","Span","Span"]
["Cusum","Cusum","Cusum"]
["kaart","kaart","chart"]
["Shewhart","Shewhart","Shewhart"]
["Gebaseerd op default waarden (niet bedoeld voor nieuwe meetsystemen)","Gebaseerd op default waarden (niet bedoeld voor nieuwe meetsystemen)","Based on default values (not meant for new AMS)"]
["Programma gegevens","Programma gegevens","Program data"]
["Correctiefactor","Correctiefactor","Correction factor"]
["Uitlezing","Uitlezing","Reading"]
["Ref","Ref","Ref"]
["Berekende bijstelling voor drift","Berekende bijstelling voor drift","Calculated adjustment for drift"]
["Bijstelling","Bijstelling","Adjustment"]
["Bijstelling voor drift noodzakelijk","Bijstelling voor drift noodzakelijk","Adjustment for drift required"]
["Drift correcties altijd uitgevoerd","Drift correcties altijd uitgevoerd","Adjustments for drift always executed"]
["Probleem met precisie; onderhoud noodzakelijk","Probleem met precisie; onderhoud noodzakelijk","Problem with precision; maintenance required"]
["Onderhoud uitgevoerd?","Onderhoud uitgevoerd?","Maintenance performed?"]
["Benodigd onderhoud altijd uitgevoerd","Benodigd onderhoud altijd uitgevoerd","Required maintenance always executed"]
["Laatste QAL3 procedure","Laatste QAL3 procedure","Last QAL3-procedure"]
["Jaar","Jaar","Year"]
["Test","Test","Test"]
["Uitlijning en vervuiling","Uitlijning en vervuiling","Alignment and cleanliness"]
["Monsternamesysteem","Monsternamesysteem","Sampling system"]
["Documentatie en verslagen","Documentatie en verslagen","Documentation and records"]
["Onderhoudgeschiktheid","Onderhoudgeschiktheid","Serviceability"]
["Controle van nul en span","Controle van nul en span","Zero and span check"]
["Lektoets","Lektoets","Leak test"]
["Lineariteit","Lineariteit","Linearity test"]
["Invloed storende factoren","Invloed storende factoren","Interferences"]
["Drift van nul en span (audit)","Drift van nul en span (audit)","Zero & span drift (audit)"]
["Responstijd","Responstijd","Response time"]
["Rapportage","Rapportage","Report"]
["Uitvoerder: naam / bedrijf","Uitvoerder: naam / bedrijf","Operator: name / firm"]
["Methode","Methode","Method"]
["Resultaat","Resultaat","Result"]
["Tests uitgevoerd","Tests uitgevoerd","Tests executed"]
["Allemaal","Allemaal","All"]
["Niet allemaal","Niet allemaal","Not all"]
["Tijd tussen 2 parallelmetingen kleiner dan 1 uur","Tijd tussen 2 parallelmetingen kleiner dan 1 uur","Time between 2 parallel measurements less than 1 hour"]
["uitgevoerd","uitgevoerd","executed"]
["Gegegevens","Gegegevens","Data"]
["Middelingstijd","Middelingstijd","Averaging time"]
["Uren buiten beschouwing (bv storing)","Uren buiten beschouwing (bv storing)","Excluded hours for evaluation (e.g. hours of disturbances)"]
["Metingen boven geldig kalibratie bereik tijdens normale bedrijfsvoering","Metingen boven geldig kalibratie bereik tijdens normale bedrijfsvoering","Measurements above valid calibration range during normal operation"]
["uur","uur","hour"]
["Periode","Periode","Period"]
["Aantal","Aantal","Number"]
["Conclusie","Conclusie","Conclusion"]
["QAL2 binnen 6 maanden","QAL2 binnen 6 maanden","QAL2 within 6 month"]
["(of AST)","(of AST)","(or AST)"]
["Onbeveiligde werkbladen","Onbeveiligde werkbladen","Unprotected sheets"]
["Laatste test","Laatste test","Last test"]
["Vervuiling optische oppervlaktes","Vervuiling optische oppervlaktes","Contamination of optical surfaces"]
["Totale onzekerheid t.g.v. selectiviteit *","Totale onzekerheid t.g.v. selectiviteit *","Total uncertainty due to selectivity *"]
["Volle schaal te laag: ","Volle schaal te laag: ","Full scale too low: "]
["Geen gegevens","Geen gegevens","No data"]
[" QAL3 van"," QAL3 van","QAL3 from"]
[" tot "," tot "," until "]

{"sheet":"Version","index":2,"n_rows":20,"n_cols":6,"rows_kept":20,"cut":false,"hidden":true,"merged":[[1,1,1,4],[2,1,2,4]],"flags":["formulas_cached_only"]}
[]
[null,"Versiebeheer"]
[null,"EN14181-NOx.xls Versie 2.1"]
[]
[null,"Versie 2.1","Datum","Te gebruiken tot","Veranderingen t.o.v. vorige versie","Test"]
[null,"1.1","29-05-07","28-06-07","werkblad version toegevoegd; test op uitvoering functionaliteitstesten QAL2 en AST aangepast; aanvinkvakje voor programma-ontwikkeling toegevoegd waarmee alle werkbladen zichtbaar worden; tevens wordt er op het werkblad menu aangegeven welke werkbladen niet beveiligd zijn; foutieve correctie voor NO2-aandeel in QAL1 aangepast; attendering voor driftproblemen nu ook voor span op QAL3-blad.","1.2"]
[null,"1.2","28-06-07","2007-09-01","Voor de test op overschrijding valid calibration range het 105% overschrijdingscriterium aangepast voor periodes tussen AST of QAL2.","1.2a"]
[null,"1.2a","05-07-07","2007-09-01","In het blad <QAL1 Selectivity> werd ten onrechte de onzekerheid t.g.v. kruisgevoeligheid eerst gedeeld door de max. concentratie van de interferent. Dit is verbeterd. De QAL1-defaults van SICK overgenomen. Vervuiling van de optische delen toegevoegd.","1.3"]
[null,"1.3","05-07-07","2007-09-01","In de bladen <QAL1 S-AMS> en < QAL1 S-AMS (zero)> wordt de drift en vervuiling als 95% BI opgegeven. Voor de berekening van de onzekerheid die als stdev wordt uitgedrukt, moesten deze waarden nog door wortel drie worden gedeeld. Dit is verbeterd.","1.3a"]
[null,"1.3a","07-07-07","2007-09-01","In blad <QAL1> de fout t.g.v. NO2-aandeel bij NO-metingen zonder NO2-metingen verandert in 20% van het NO2-aandeel. In blad <VCR-test> de formules aangepast; de teller voor 5% overschrijding was onjuist.","1.4"]
[null,"1.4","08-07-07","2007-09-01","In blad <QAL1> het NO2-aandeel bij NO-metingen zonder NO2-metingen als NO2-meting doorgevoerd met een stdev van 15%; door deze wijziging komt de herleide NOx-concentartie in blad <QAL1 numdif> altjd overeen met de ingevoerde concentratie in <QAL1>. ","1.5"]
[null,"1.5","20-09-07","2008-01-01","ABB gegevens verwerkt; eenheden <QAL1 Selectivity> onjuist, aangepast in %TV; default selectiviteit als een totale onzekerheid ingebracht","1.6"]
[null,"1.6","02-10-07","2011-01-01","Default onzekerheid voor monsternameverliezen voor extractieve NO2-meting op 5% gesteld.","1.7"]
[null,"1.7","29-10-07","2011-01-01","Fout uit versiebeheer er uitgehaald. Variabilieitstoets en geldigheid kalibratiecurvetoets in QAL2 en AST wordt nu uitgevoerd mbv de standaard deviatie. In de vorige versie ging dit fout en werden de toetsen op basis van de onzekerheid uitgevoerd.","2.1"]
[null,"1.8","15-4-08","2011-01-01","Eenheid NO-, NO2- en Nox kan gespecificeerd worden. Emissies kunnen ook in g/GJ worden uitgelezen. Actuele rookgasdruk wordt nu ook in berekeningen meegenomen. De cusum-kaart wordt nu ook gereset bij onderhoud.","1.8"]
[null,"1.9","16-4-08","2011-01-01"," In de AST is een factor 1,5 in de variabiliteitstoets toegevoegd en overig commentaar Kema verwerkt.","1.9"]
[null,"2,0","11-7-08","2011-01-01","bij de functionaliteitstoetsen gaf tool ten onrechte aan dat alle test waren uitgevoerd, ook wanneer er slechts 1 kolom werd ingevuld","2,0"]
[null,"2.0","14-7-08","2011-01-01","Bij de AST metingen vond de toetsing op data niet correct plaats. De cellen waren 'verkeerd' gevalideerd.","2.0"]
[null,"2.1","14-7-08","2015-01-01","fout gecorrigeerd in QAL3 sheet, laatste QAL3 uitvoering altijd Nox datum. Verder tabblad menu aangepast QAL3 begin en einddatum nu weergegeven wanneer er geen problemen zijn","2.1"]
[{"$nulls":5},"2.1"]

{"sheet":"Menu","index":3,"n_rows":27,"n_cols":8,"rows_kept":27,"cut":false,"hidden":false,"merged":[[26,1,26,8],[1,1,1,8],[2,1,2,8],[3,1,3,2],[22,2,22,3]],"flags":["formulas_cached_only"]}
[]
[null,"NEN-EN 14181 voor NOx-meetsystemen"]
[null,"EN14181-NOx.xls Versie 2.1"]
[null,"Taal"]
[null,null,"QAL1","Ontbrekende gegevens concentratieprofiel","Berekend betrouwbaarheidsinterval",""," %"]
[null,null,null,"S-AMS ","Gebaseerd op default waarden (niet bedoeld voor nieuwe meetsystemen)"]
[null,null,null,"Selectiviteit","Gebaseerd op default waarden (niet bedoeld voor nieuwe meetsystemen)"]
[null,null,null,"Overige onzekerheidsbronnen","Gebaseerd op default waarden (niet bedoeld voor nieuwe meetsystemen)"]
[null,null,null,"Numerieke differentiatie"]
[null,null,null,"S-AMS nul ","Gebaseerd op default waarden (niet bedoeld voor nieuwe meetsystemen)"]
[null,null,"QAL2","(Jaar 0: 2015)",""]
[null,null,null,"Kalibratie",""]
[null,null,null,"Functionele testen","Niet allemaal uitgevoerd"]
[null,null,null,"Wekelijkse controle geldig kalibratiebereik (VCR-test)",""]
[null,null,"QAL3"]
[null,null,null,"Cusum NO-meting","Geen gegevens","","",""]
[null,null,null,"Cusum NO2-meting","Geen gegevens","","",""]
[null,null,null,"Cusum O2-meting","Geen gegevens","","",""]
[null,null,null,"","","","",""]
[null,null,"AST"]
[null,null,null,"Parallelmetingen","Kalibratie lijn niet-valide"]
[null,null,null,"Functionele testen","Niet allemaal uitgevoerd"]
[null,null,"Programma ontwikkeling"]
[null,null,null,"Versiebeheer"]
[null,null,null,"Taal"]
[null,null,null,"Programma gegevens"]
[null,"Ondanks de zorgvuldigheid waarmee dit werkblad is opgesteld, zijn fouten niet uit te sluiten."]

{"sheet":"QAL1","index":4,"n_rows":63,"n_cols":9,"rows_kept":63,"cut":false,"hidden":true,"merged":[[60,1,60,8],[1,1,1,8],[2,1,2,8],[56,1,56,8],[59,1,59,8]],"flags":["formulas_cached_only"]}
[]
[null,"Resultaten QAL1 voor NOx-meetsystemen"]
[null,"EN14181-NOx.xls Versie 2.1"]
[]
[null,"Specificatie meetsysteem"]
[null,"NOx-meting"]
[null,"O2-meting"]
[null,"Vochtmeting"]
[null,"Rookgastemperatuurmeting"]
[null,"Rookgasdrukmeting"]
[null,"Meting","NOx","NO","NO2","O2","H2O","Rookgas-temperatuur","Absolute rookgas druk"]
[null,"Invoer noodzakelijk","Nee","Ja","Ja","Ja","Nee","Ja","Ja"]
[null,"Volle schaal"]
[null,"Eenheid","","vppm","vppm","vol%",""]
[null,null,"","vppm","vppm"]
[null,"Onzekerheid QAL3-medium (%) *"]
[null,null,"",0,0]
[null,"NOx-concentratie"]
[null,"Referentie zuurstofconcentratie",null,"vol% droog rookgas"]
[null,"Emissiegrenswaarde",null,"mg/Nm3 (als NO2) droog rookgas bij  vol% O2",{"$nulls":4},""]
[null,"Testwaarde bij standaard condities",null,"mg/Nm3 (als NO2) droog rookgas bij  vol% O2",{"$nulls":4},""]
[null,"Vereist betrouwbaarheidsinterval",null,"%"]
[]
[null,"Rookgascondities"]
[null,"Geschatte O2-concentratie",null,"vol% droog rookgas"]
[null,"Geschat aandeel NO2",null,"%"]
[null,"Geschatte H20-concentratie",null,"vol% nat rookgas"]
[null,"Geschatte rookgastemperatuur",null,"°C"]
[null,"Geschatte absolute druk van het rookgas",null,"hPa"]
[null,5,"Referentie zuurstofconcentratie Emissiegrenswaarde Testwaarde bij standaard condities Vereist betrouwbaarheidsinterval","O2 NO2 H2O Rookgas-temperatuur Absolute rookgas druk"," NO NO2 O2 "," NO NO2 O2 "]
[null,"Berekening onzekerheid meetinstrument","NOx","NO","NO2","O2","H2O","Rookgas-temperatuur","Absolute rookgas druk"]
[null,"Volle schaal","",0,0,0,0]
[null,"Testwaarde","",0,0,"",0,0,1013]
[null,"Eenheid","","mg/Nm3","mg/Nm3","vol%","vol%","°C","hPa"]
[null,null,"","als NO2","als NO2"]
[null,"S-AMS (%) *","",0.4655188815848104,0.4655188815848104,0.3705084631288718,"","",""]
[null,"Onzekerheid t.g.v. selectiviteit (%) *","",5.2,5.2,0.16124515496597103,"","",""]
[null,"Onzekerheid t.g.v. overige bronnen (%) *","",1.4457332071211617,3.2285411317662,0.3036410984479319,"","",""]
[null,"Onzekerheid t.g.v. converterrendement (%) *",""]
[null,"Onzekerheid in aanname NO2-aandeel (%) *",null,null,""]
[null,"Totale onzekerheid meting (%) *","",5.417273514904413,6.138418824796672,0.5054447921551722,2,1,2]
[]
[null,"Resultaten concentratieprofielmeting"]
[null,"Hoogste herleide NOx-concentratie",null," mg/Nm3 (als NO2) droog rookgas bij  vol% O2"]
[null,"Laagste herleide NOx-concentratie",null," mg/Nm3 (als NO2) droog rookgas bij  vol% O2"]
[null,"Aantal meetpunten AMS"]
[null,"Geschatte onzekerheid t.g.v. niet representatieve monstername AMS *",""," %"]
[]
[null,"Betrouwbaarheidsinterval"]
[null,"Meetonzekerheid herleide NOx-concentratie *",""," %"]
[null,"Berekend betrouwbaarheidsinterval",""," %"]
[null,"Vereist betrouwbaarheidsinterval",20," %"]
[null,"Resultaat toetsing betrouwbaarheidsinterval",""]
[]
[null,"OPMERKINGEN"]
[null,"Voer data in voor: Referentie zuurstofconcentratie Emissiegrenswaarde Testwaarde bij standaard condities Vereist betrouwbaarheidsinterval"]
[null,"Voer geschatte waarden in voor:  O2 NO2 H2O Rookgas-temperatuur Absolute rookgas druk"]
[null,""]
[null,"Ontbrekende gegevens concentratieprofiel"]
[null,"Voer volle schaal in voor:   NO NO2 O2 "]
[null,"Voer onzekerheid QAL3-medium in voor:   NO NO2 O2 "]
[]
[null," *: als standaard deviatie"]

{"sheet":"QAL1 S-AMS","index":5,"n_rows":37,"n_cols":11,"rows_kept":37,"cut":false,"hidden":true,"merged":[[1,1,1,10],[2,1,2,10]],"flags":["formulas_cached_only"]}
[]
[null,"Berekening S-AMS van het NOx-meetsysteem"]
[null,"EN14181-NOx.xls Versie 2.1"]
[]
[null,"Default waarden",null,"NOx","NO","NO2","O2","H2O","Rookgas-temperatuur","Absolute rookgas druk","Eenheid"]
[null,"Waarden als % van de volle schaal",null,true,true,true,true,true]
[null,"Instabiliteit (drift)",null,0.44490909353363,0.44490909353363,0.44490909353363,0.47874836814343297,0.5611260999098153,0,0,"%"]
[null,"Vervuiling optische oppervlaktes",null,-0.5,-0.5,-0.5,0,4.8,null,null,"%"]
[null,"Gevoeligheid voor omgevingsdruk",null,0,0,0,0,0,0,0,"%/kPa"]
[null,"Gevoeligheid voor omgevingstemperatuur",null,0.010022398144531304,0.010022398144531304,0.010022398144531304,0.015217549750066147,0.09152908729139463,0,0,"%/K"]
[null,"Gevoeligheid voor variaties netspanning",null,0.13416407864998736,0.13416407864998736,0.13416407864998736,0.10000001977336484,0.27,0,0,"%/10 V"]
[]
[null,"Instrument specifieke waarden",null,"NOx","NO","NO2","O2","H2O","Rookgas-temperatuur","Absolute rookgas druk","Eenheid"]
[null,"Waarden als % van de volle schaal",null,false,false,false,false,false]
[null,"Instabiliteit (drift)",{"$nulls":8},"%"]
[null,"Vervuiling optische oppervlaktes",{"$nulls":8},"%"]
[null,"Gevoeligheid voor omgevingsdruk",{"$nulls":8},"%/kPa"]
[null,"Gevoeligheid voor omgevingstemperatuur",{"$nulls":8},"%/K"]
[null,"Gevoeligheid voor variaties netspanning",{"$nulls":8},"%/10 V"]
[{"$nulls":10},"%"]
[null,"Waarden als percentage van de volle schaal",null,true,true,true,true,true]
[null,"Berekening S-AMS","Variaties","NOx","NO","NO2","O2","H2O","Rookgas-temperatuur","Absolute rookgas druk","Eenheid"]
[null,"Instabiliteit (drift) *",null,0.25686838491655367,0.25686838491655367,0.25686838491655367,0.2764054992217051,0.32396630483225675,0,0,"% TW"]
[null,"Vervuiling optische oppervlaktes *",null,-0.2886751345948129,-0.2886751345948129,-0.2886751345948129,0,2.7712812921102037,0,0,"% TW"]
[null,"Gevoeligheid voor omgevingsdruk",null,0,0,0,0,0,0,0,"%/kPa"]
[null,null,5,{"$nulls":7},"kPa"]
[null,"Onzekerheid *",null,0,0,0,0,0,0,0,"% TW"]
[null,"Gevoeligheid voor omgevingstemperatuur",null,0.010022398144531304,0.010022398144531304,0.010022398144531304,0.015217549750066147,0.09152908729139463,0,0,"%/K"]
[null,null,20,{"$nulls":7},"K"]
[null,"Onzekerheid *",null,0.11572868533341509,0.11572868533341509,0.11572868533341509,0.17571712889214428,1.056886863727349,0,0,"% TW"]
[null,"Gevoeligheid voor variaties netspanning",null,0.13416407864998736,0.13416407864998736,0.13416407864998736,0.10000001977336484,0.27,0,0,"%/10 V"]
[null,null,30,{"$nulls":7},"V"]
[null,"Onzekerheid *",null,0.23237900077244503,0.23237900077244503,0.23237900077244503,0.1732051150053603,0.467653718043597,0,0,"% TW"]
[null,"",null,0,0,0,0,0,0,0,"% TW"]
[null,"S-AMS *",null,"",0.4655188815848104,0.4655188815848104,0.3705084631288718,"","","","% TW"]
[]
[null," *: als standaard deviatie",{"$nulls":8},"% TW: % testwaarde"]

{"sheet":"QAL1 Selectivity","index":6,"n_rows":55,"n_cols":13,"rows_kept":55,"cut":false,"hidden":true,"merged":[[4,1,5,1],[20,1,21,1],[20,2,20,4],[20,5,20,12],[37,1,38,1],[1,1,1,12],[2,1,2,12],[37,2,37,4],[37,5,37,12],[4,5,4,12],[4,2,4,4]],"flags":["formulas_cached_only"]}
[]
[null,"Berekening selectiviteit van het NOx-meetsysteem"]
[null,"EN14181-NOx.xls Versie 2.1"]
[]
[null,"Component","Variaties rookgassamenstelling",null,null,"Default selectiviteit"]
[null,null,"Min","Max","Eenheid","NOx","NO","NO2","O2","H2O","Rookgas-temperatuur","Absolute rookgas druk","Eenheid"]
[null,"Waarden als % van de volle schaal",null,null,null,false,false,false,false,false]
[null,"O2",3,21,"vol%",0,0,0,0,0,0,0,"% / vol%"]
[null,"H2O",1,30,"vol%",0,0,0,0,0,0,0,"% / vol%"]
[null,"CO",0,50,"ppm",0,0,0,0,0,0,0,"% / ppm"]
[null,"CO2",0,15,"vol%",0,0,0,0,0,0,0,"% / vol%"]
[null,"CH4",0,50,"ppm",0,0,0,0,0,0,0,"% / ppm"]
[null,"N2O",0,20,"ppm",0,0,0,0,0,0,0,"% / ppm"]
[null,"NO ",0,300,"ppm",0,0,0,0,0,0,0,"% / ppm"]
[null,"NO2",0,30,"ppm",0,0,0,0,0,0,0,"% / ppm"]
[null,"NH3",0,20,"ppm",0,0,0,0,0,0,0,"% / ppm"]
[null,"SO2",0,200,"ppm",0,0,0,0,0,0,0,"% / ppm"]
[null,"HCl",0,50,"ppm",0,0,0,0,0,0,0,"% / ppm"]
[null,"Totale onzekerheid t.g.v. selectiviteit *",null,null,null,5.2,5.2,5.2,0.16124515496597103,0.21230229219680158,null,null,"% TW"]
[]
[null,"Component","Variaties rookgassamenstelling",null,null,"Instrument specifieke selectiviteit"]
[null,null,"Min","Max","Eenheid","NOx","NO","NO2","O2","H2O","Rookgas-temperatuur","Absolute rookgas druk","Eenheid"]
[null,"Waarden als % van de volle schaal",null,null,null,false,false,false,false,false]
[null,"O2",null,null,"vol%",{"$nulls":7},"% / vol%"]
[null,"H2O",null,null,"vol%",{"$nulls":7},"% / vol%"]
[null,"CO",null,null,"ppm",{"$nulls":7},"% / ppm"]
[null,"CO2",null,null,"vol%",{"$nulls":7},"% / vol%"]
[null,"CH4",null,null,"ppm",{"$nulls":7},"% / ppm"]
[null,"N2O",null,null,"ppm",{"$nulls":7},"% / ppm"]
[null,"NO ",null,null,"ppm",{"$nulls":7},"% / ppm"]
[null,"NO2",null,null,"ppm",{"$nulls":7},"% / ppm"]
[null,"NH3",null,null,"ppm",{"$nulls":7},"% / ppm"]
[null,"SO2",null,null,"ppm",{"$nulls":7},"% / ppm"]
[null,"HCl",null,null,"ppm",{"$nulls":7},"% / ppm"]
[{"$nulls":12},""]
[{"$nulls":12},""]
[]
[null,"Component","Variaties rookgassamenstelling",null,null,"Selectiviteit"]
[null,null,"Min","Max","Eenheid","NOx","NO","NO2","O2","H2O","Rookgas-temperatuur","Absolute rookgas druk","Eenheid"]
[null,"O2",3,21,"vol%",0,0,0,0,0,0,0,"% TW"]
[null,"H2O",1,30,"vol%",0,0,0,0,0,0,0,"% TW"]
[null,"CO",0,50,"ppm",0,0,0,0,0,0,0,"% TW"]
[null,"CO2",0,15,"vol%",0,0,0,0,0,0,0,"% TW"]
[null,"CH4",0,50,"ppm",0,0,0,0,0,0,0,"% TW"]
[null,"N2O",0,20,"ppm",0,0,0,0,0,0,0,"% TW"]
[null,"NO ",0,300,"ppm",0,0,0,0,0,0,0,"% TW"]
[null,"NO2",0,30,"ppm",0,0,0,0,0,0,0,"% TW"]
[null,"NH3",0,20,"ppm",0,0,0,0,0,0,0,"% TW"]
[null,"SO2",0,200,"ppm",0,0,0,0,0,0,0,"% TW"]
[null,"HCl",0,50,"ppm",0,0,0,0,0,0,0,"% TW"]
[null,"","","","",0,0,0,0,0,0,0,""]
[null,"","","","",0,0,0,0,0,0,0,""]
[null,"Totale onzekerheid t.g.v. selectiviteit *",null,null,null,"",5.2,5.2,0.16124515496597103,0.21230229219680158,0,0,"% TW"]
[]
[null," *: als standaard deviatie",{"$nulls":10},"% TW: % testwaarde"]

{"sheet":"QAL1 other sources","index":7,"n_rows":27,"n_cols":11,"rows_kept":27,"cut":false,"hidden":true,"merged":[[1,1,1,10],[2,1,2,10]],"flags":["formulas_cached_only"]}
[]
[null,"Berekening onzekerheid van het NOx-meetsysteem t.g.v. overig bronnen "]
[null,"EN14181-NOx.xls Versie 2.1"]
[]
[null,"Default waarden",null,"NOx","NO","NO2","O2","H2O","Rookgas-temperatuur","Absolute rookgas druk","Eenheid"]
[null,"Afwijking lineariteit *",null,0.697853852894716,0.697853852894716,0.697853852894716,0.3663673020344474,1.2895008724308796,0,0,"% VS"]
[null,"Herhaalbaarheid *",null,1.388456399329668,1.388456399329668,1.388456399329668,0.2178445546714446,0.5892851601728996,0,0,"% TW"]
[null,"Verliezen in monsternamesysteem *",null,0,0,5,0,0,0,0,"% TW"]
[null,null,null,0,0,0,0,0,0,0,"% TW"]
[]
[null,"Instrument specifieke waarden",null,"NOx","NO","NO2","O2","H2O","Rookgas-temperatuur","Absolute rookgas druk","Eenheid"]
[null,"Afwijking lineariteit *",{"$nulls":8},"% VS"]
[null,"Herhaalbaarheid *",{"$nulls":8},"% TW"]
[null,"Verliezen in monsternamesysteem *",{"$nulls":8},"% TW"]
[null,"",{"$nulls":8},"% TW"]
[{"$nulls":10},"% TW"]
[null,"Waarden als percentage van de volle schaal"]
[null,"Overige onzekerheidsbronnen",null,"NOx","NO","NO2","O2","H2O","Rookgas-temperatuur","Absolute rookgas druk","Eenheid"]
[null,"Afwijking lineariteit *",null,0.40290610982378183,0.40290610982378183,0.40290610982378183,0.21152226045186515,0.7444936758182257,0,0,"% TW"]
[null,"Herhaalbaarheid *",null,1.388456399329668,1.388456399329668,1.388456399329668,0.2178445546714446,0.5892851601728996,0,0,"% TW"]
[null,"Verliezen in monsternamesysteem *",null,0,0,2.886751345948129,0,0,0,0,"% TW"]
[null,"",null,0,0,0,0,0,0,0,"% TW"]
[null,"",null,0,0,0,0,0,0,0,"% TW"]
[null,"Onzekerheid t.g.v. overige bronnen *",null,"",1.4457332071211617,3.2285411317662,0.3036410984479319,0.9494881954681339,0,0,"% TW"]
[]
[null," *: als standaard deviatie",{"$nulls":8},"% VS: % volle schaal"]
[{"$nulls":10},"% TW: % testwaarde"]

{"sheet":"QAL1 NumDif","index":8,"n_rows":37,"n_cols":32,"rows_kept":37,"cut":false,"hidden":true,"merged":[[5,4,5,5],[4,1,5,3],[1,1,1,16],[2,1,2,16],[3,4,3,5],[4,4,4,5]],"flags":["formulas_cached_only"]}
[]
[null,"Berekening van de totale onzekerheid m.b.v. numerieke differentiatie"]
[null,"EN14181-NOx.xls Versie 2.1"]
[{"$nulls":4},"regelnummer",null,5,6,7,8,9,10,11,12,13,14,15,16,17,18,19,20,21,22,23,24,25,26,27,28,29]
[null,"Spreadsheet methode",null,null,"Parameter",null,"NOx","NO","NO2","T NOx","P NOx","Vocht NOx","O2","T O2","P O2","Vocht O2","Representativiteit","","","","","","","","","","","","","",""]
[{"$nulls":4},"Δ",null,0,0,0,0,0,0,{"$error":"#VALUE!"},0,20.26,0,{"$error":"#VALUE!"},"","","","","","","","","","","","","",""]
["kolom","Parameter","Waarde","Eenheid","St. dev.","Relatief?"]
["G","NOx",0,"",0,"Ja",0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0]
["H","NO",0,"mg/Nm3",5.417273514904413,"Ja",0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0]
["I","NO2",0,"mg/Nm3",6.138418824796672,"Ja",0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0]
["J","T NOx",0,"°C",0,"Ja",0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0]
["K","P NOx",1013,"hPa",0,"Ja",1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013]
["L","Vocht NOx",0,"vol%",0,"Ja",0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0]
["M","O2","","vol%",0.5054447921551722,"Ja","","","","","","","","","","","","","","","","","","","","","","","","","",""]
["N","T O2",0,"°C",1,"Ja",0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0]
["O","P O2",1013,"hPa",2,"Ja",1013,1013,1013,1013,1013,1013,1013,1013,1033.26,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013,1013]
["P","Vocht O2",0,"vol%",2,"Ja",0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0,0]
["Q","Representativiteit",1,null,"","Ja",1,1,1,1,1,1,1,1,1,1,{"$error":"#VALUE!"},1,1,1,1,1,1,1,1,1,1,1,1,1,1,1]
["R",{"$nulls":4},"Ja","","","","","","","","","","","","","","","","","","","","","","","","","",""]
["S",{"$nulls":4},"Ja","","","","","","","","","","","","","","","","","","","","","","","","","",""]
["T",{"$nulls":4},"Ja","","","","","","","","","","","","","","","","","","","","","","","","","",""]
["U",{"$nulls":4},"Ja","","","","","","","","","","","","","","","","","","","","","","","","","",""]
["V",{"$nulls":4},"Ja","","","","","","","","","","","","","","","","","","","","","","","","","",""]
["W",{"$nulls":4},"Ja","","","","","","","","","","","","","","","","","","","","","","","","","",""]
["X",{"$nulls":4},"Ja","","","","","","","","","","","","","","","","","","","","","","","","","",""]
["Y",{"$nulls":4},"Ja","","","","","","","","","","","","","","","","","","","","","","","","","",""]
["Z",{"$nulls":4},"Ja","","","","","","","","","","","","","","","","","","","","","","","","","",""]
["AA",{"$nulls":4},"Ja","","","","","","","","","","","","","","","","","","","","","","","","","",""]
["AB",{"$nulls":4},"Ja","","","","","","","","","","","","","","","","","","","","","","","","","",""]
["AC",{"$nulls":4},"Ja","","","","","","","","","","","","","","","","","","","","","","","","","",""]
["AD",{"$nulls":4},"Ja","","","","","","","","","","","","","","","","","","","","","","","","","",""]
["AE",{"$nulls":4},"Ja","","","","","","","","","","","","","","","","","","","","","","","","","",""]
[null,"NOx-concentratie bij standaard condities",{"$error":"#VALUE!"},"mg/Nm3 (als NO2) droog rookgas bij  vol% O2",{"$error":"#VALUE!"},"Ja",{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"}]
[{"$nulls":4},{"$error":"#VALUE!"},{"$error":"#VALUE!"}]
[null,"Δ",{"$nulls":4},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"}]
[null,"Δ2",{"$nulls":4},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"},{"$error":"#VALUE!"}]
[null,"Σ(Δ2)",{"$error":"#VALUE!"}]

{"sheet":"QAL1 S-AMS (zero)","index":9,"n_rows":37,"n_cols":11,"rows_kept":37,"cut":false,"hidden":true,"merged":[[1,1,1,10],[2,1,2,10]],"flags":["formulas_cached_only"]}
[]
[null,"Berekening S-AMS (zero) van het NOx-meetsysteem"]
[null,"EN14181-NOx.xls Versie 2.1"]
[]
[null,"Default waarden",null,"NOx","NO","NO2","O2","H2O","Rookgas-temperatuur","Absolute rookgas druk","Eenheid"]
[null,"Waarden als % van de volle schaal",null,true,true,true,true,true]
[null,"Instabiliteit (drift)",null,0.6358616201658974,0.6358616201658974,0.6358616201658974,0.47874836814343297,0.5611260999098153,0,0,"%"]
[null,"Vervuiling optische oppervlaktes",null,-0.5,-0.5,-0.5,0,4.8,null,null,"%"]
[null,"Gevoeligheid voor omgevingsdruk",null,0,0,0,0,0,0,0,"%/kPa"]
[null,"Gevoeligheid voor omgevingstemperatuur",null,0.09036038955205981,0.09036038955205981,0.09036038955205981,0.015217549750066147,0.09152908729139463,0,0,"%/K"]
[null,"Gevoeligheid voor variaties netspanning",null,0.17464249196572978,0.17464249196572978,0.17464249196572978,0.10000001977336484,0.27,0,0,"%/10 V"]
[]
[null,"Instrument specifieke waarden",null,"NOx","NO","NO2","O2","H2O","Rookgas-temperatuur","Absolute rookgas druk","Eenheid"]
[null,"Waarden als % van de volle schaal",null,false,false,false,false,false]
[null,"Instabiliteit (drift)",{"$nulls":8},"%"]
[null,"Vervuiling optische oppervlaktes",{"$nulls":8},"%"]
[null,"Gevoeligheid voor omgevingsdruk",{"$nulls":8},"%/kPa"]
[null,"Gevoeligheid voor omgevingstemperatuur",{"$nulls":8},"%/K"]
[null,"Gevoeligheid voor variaties netspanning",{"$nulls":8},"%/10 V"]
[{"$nulls":10},"%"]
[null,"Waarden als percentage van de volle schaal",null,true,true,true,true,true]
[null,"Berekening S-AMS (nul)","Variaties","NOx","NO","NO2","O2","H2O","Rookgas-temperatuur","Absolute rookgas druk","Eenheid"]
[null,"Instabiliteit (drift) *",null,0.36711487757013245,0.36711487757013245,0.36711487757013245,0.2764054992217051,0.32396630483225675,0,0,"% TW"]
[null,"Vervuiling optische oppervlaktes *",null,-0.2886751345948129,-0.2886751345948129,-0.2886751345948129,0,2.7712812921102037,0,0,"% TW"]
[null,"Gevoeligheid voor omgevingsdruk",null,0,0,0,0,0,0,0,"%/kPa"]
[null,null,5,{"$nulls":7},"kPa"]
[null,"Onzekerheid *",null,0,0,0,0,0,0,0,"% TW"]
[null,"Gevoeligheid voor omgevingstemperatuur",null,0.09036038955205981,0.09036038955205981,0.09036038955205981,0.015217549750066147,0.09152908729139463,0,0,"%/K"]
[null,null,20,{"$nulls":7},"K"]
[null,"Onzekerheid *",null,1.0433919046392237,1.0433919046392237,1.0433919046392237,0.17571712889214428,1.056886863727349,0,0,"% TW"]
[null,"Gevoeligheid voor variaties netspanning",null,0.17464249196572978,0.17464249196572978,0.17464249196572978,0.10000001977336484,0.27,0,0,"%/10 V"]
[null,null,20,{"$nulls":7},"V"]
[null,"Onzekerheid *",null,0.2016597794967223,0.2016597794967223,0.2016597794967223,0.11547007667024019,0.31176914536239797,0,0,"% TW"]
[null,"",null,0,0,0,0,0,0,0,"% TW"]
[null,"S-AMS nul *",null,"",1.160792832507162,1.160792832507162,0.3472892857436429,"","","","% TW"]
[]
[null," *: als standaard deviatie",{"$nulls":8},"% TW: % testwaarde"]

{"sheet":"QAL2","index":10,"n_rows":18,"n_cols":5,"rows_kept":18,"cut":false,"hidden":true,"merged":[[1,1,1,6],[2,1,2,6]],"flags":["formulas_cached_only"]}
[]
[null,"Resultaten QAL2 voor NOx-meetsystemen"]
[null,"EN14181-NOx.xls Versie 2.1"]
[]
[null,"Kalibratie"]
[null,"Datum","2015-01-01"," - ","2015-01-01"]
[null,"Meetinstantie",""]
[null,"Aantal parallelmetingen",""]
[null,"Nieuwe kalibratie lijn"]
[null,"Intercept",""]
[null,"Helling",""]
[null,"Geldig kalibratiebereik","","mg/Nm3 (als NO2) droog rookgas bij  vol% O2"]
[null,"Variabiliteitstest",{"$error":"#NUM!"}]
[null,"Opmerkingen","Te weinig metingen!"]
[]
[null,"Functionele testen"]
[null,"Datum","","",""]
[null,"Tests uitgevoerd","Niet allemaal"]

{"sheet":"QAL2 calibration","index":11,"n_rows":71,"n_cols":18,"rows_kept":71,"cut":false,"hidden":true,"merged":[[23,13,24,13],[23,17,24,17],[23,14,24,14],[23,12,24,12],[23,15,24,15],[23,16,24,16],[22,4,22,10],[23,10,24,10],[17,6,17,10],[18,6,18,10],[22,1,23,2],[24,2,24,3],[23,4,23,6],[23,7,23,9],[15,6,15,10],[16,6,16,10],[1,1,1,10],[2,1,2,10],[5,6,5,10],[6,6,6,10]],"flags":["formulas_cached_only"]}
[]
[null,"Kalibratie van het NOx-meetsysteem"]
[null,"EN14181-NOx.xls Versie 2.1"]
[]
[null,"Specificatie meetsysteem"]
[null,"NOx-meetmethode",{"$nulls":4},"Extractieve NO- + NO2-meting in droog rookgas"]
[null,"O2-meetmethode",{"$nulls":4},"In-situ O2-meting"]
[null,"Kalibratielijn gebruikt tijdens kalibratie"]
[null,null,"Intercept"]
[null,null,"Helling"]
[null,"Emissiegrenswaarde",{"$nulls":4},0,"mg/Nm3 (als NO2) droog rookgas bij  vol% O2"]
[null,"Testwaarde bij standaard condities",{"$nulls":4},0,"mg/Nm3 (als NO2) droog rookgas bij  vol% O2"]
[null,"Vereist betrouwbaarheidsinterval",{"$nulls":4},0,"%"]
[]
[null,"Standaard referentie methode (SRM)"]
[null,"Meetinstantie"]
[null,"NOx-meetmethode"]
[null,"O2-meetmethode"]
[null,"Meetplaats"]
[null,"Monsternamemethode"]
[null,"Aantal meetpunten SRM"]
[]
[null,"Kalibratiepunten",null,null,"NOx-concentratie bij standaard condities (mg/Nm3 (als NO2) droog rookgas bij  vol% O2)"]
[{"$nulls":4},"Standaard referentie methode (SRM)",null,null,"NOx-meetsysteem (AMS)",null,null,"Regressie-lijn",null,"AMS","SRM","Regressie-lijn","Δ","(Δ-Δgem)2","Dag"]
[null,"Nummer","Datum",null,"Starttijd","Eindtijd","Meetwaarde","Starttijd","Eindtijd","Meetwaarde"]
[null,"","2015-01-01","112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"",null,"112015",{"$nulls":6},"",null,"","","","","",1]
[null,"Min","",null,null,null,"",null,null,"",0,{"$nulls":4},"",""]
[null,"Max","",null,null,null,"",null,null,"",0,{"$nulls":4},"","",1]
[null,"Gemiddeld",{"$nulls":4},"",null,null,"",{"$error":"#DIV/0!"},{"$nulls":4},"",""]
[]
[null,"Regressie-lijn"]
[null,"Aantal parallelmetingen",null,null,null,0,"Te weinig metingen!"]
[null,null,"Intercept",null,0,"",{"$error":"#VALUE!"}]
[null,null,"Helling",null,null,""]
[null,"Variabiliteitstest"]
[null,null,"S0",null,null,0]
[null,null,"KV",null,null,{"$error":"#NUM!"}]
[null,null,"KV x S0 ",null,null,{"$error":"#NUM!"}]
[null,null,"SD",null,null,""]
[null,null,"KV x S0 ",{"$error":"#NUM!"},"SD",{"$error":"#NUM!"}]
[]
[]
[]
[null,"Nieuwe kalibratie lijn"]
[null,null,"Intercept","(a)",null,""]
[null,null,"Helling","(b)",null,""]
[null,null,"Geldig kalibratiebereik",null,null,"","mg/Nm3 (als NO2) droog rookgas bij  vol% O2"]

{"sheet":"QAL2 functionality","index":12,"n_rows":20,"n_cols":17,"rows_kept":20,"cut":false,"hidden":true,"merged":[[1,1,1,5],[2,1,2,5]],"flags":["formulas_cached_only"]}
[]
[null,"QAL2 - functionele testen"]
[null,"EN14181-NOx.xls Versie 2.1"]
[]
[null,"Test","Datum","Uitvoerder: naam / bedrijf","Methode","Resultaat",0,"check","dat","uitv","meth","res",null,"tot","ok1","ok2","ok tot"]
[null,"Uitlijning en vervuiling",{"$nulls":4},1,1,0,0,0,0,null,0,0,0,0]
[null,"Monsternamesysteem",{"$nulls":4},2,1,0,0,0,0,null,0,0,0,0]
[null,"Documentatie en verslagen",{"$nulls":4},3,1,0,0,0,0,null,0,0,0,0]
[null,"Onderhoudgeschiktheid",{"$nulls":4},4,1,0,0,0,0,null,0,0,0,0]
[null,"Controle van nul en span",{"$nulls":4},5,1,0,0,0,0,null,0,0,0,0]
[null,"Lektoets",{"$nulls":4},6,1,0,0,0,0,null,0,0,0,0]
[null,"Lineariteit",{"$nulls":4},7,1,0,0,0,0,null,0,0,0,0]
[null,"Invloed storende factoren",{"$nulls":4},8,1,0,0,0,0,null,0,0,0,0]
[null,"Drift van nul en span (audit)",{"$nulls":4},9,1,0,0,0,0,null,0,0,0,0]
[null,"Responstijd",{"$nulls":4},10,1,0,0,0,0,null,0,0,0,0]
[null,"Rapportage",{"$nulls":4},11,1,0,0,0,0,null,0,0,0,0]
[]
[]
[]
[{"$nulls":16},0]

{"sheet":"VCR-test","index":13,"n_rows":170,"n_cols":13,"rows_kept":170,"cut":false,"hidden":true,"merged":[[8,10,9,10],[8,11,9,11],[1,1,1,8],[2,1,2,8],[8,1,9,1],[8,2,9,4],[8,5,9,5],[8,6,8,7],[8,8,9,8]],"flags":["formulas_cached_only"]}
[]
[null,"Wekelijkse controle geldig kalibratiebereik"]
[null,"EN14181-NOx.xls Versie 2.1"]
[]
[null,"Gegegevens"]
[null,"Middelingstijd"]
[null,"Geldig kalibratiebereik"," mg/Nm3 (als NO2) droog rookgas bij  vol% O2"]
[]
[null,"Datum","Periode",null,null,"Uren buiten beschouwing (bv storing)","Metingen boven geldig kalibratie bereik tijdens normale bedrijfsvoering",null,"Conclusie",null,"Periode","Periode > 105%"]
[{"$nulls":6},"Aantal","%"]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1,1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]
[null,null,"","","",null,null,"","",null,"",1]

{"sheet":"QAL3","index":14,"n_rows":25,"n_cols":9,"rows_kept":25,"cut":false,"hidden":true,"merged":[[1,1,1,8],[2,1,2,8]],"flags":["formulas_cached_only"]}
[]
[null,"Resultaten QAL3-procedure van het NOx-meetsysteem"]
[null,"EN14181-NOx.xls Versie 2.1"]
[]
[null,"Specificatie meetsysteem"]
[null,"NOx-meting","Extractieve NO- + NO2-meting in droog rookgas"]
[null,"O2-meting","In-situ O2-meting"]
[null,"Vochtmeting","Berekening op T-scrubber"]
[null,"Rookgastemperatuurmeting","Thermokopppel"]
[null,"Rookgasdrukmeting","Druksensor"]
[]
[null,"Meting","NOx","NO","NO2","O2","H2O","Rookgas-temperatuur","Absolute rookgas druk"]
[null,"Eenheid","","mg/Nm3","mg/Nm3","vol%","vol%","°C","hPa"]
[null,"Volle schaal","",0,0,0,0,"",""]
[null,"Laatste QAL3 procedure","","","","","2011-01-01"]
[null,"S-AMS nul *","",0,0,{"$error":"#VALUE!"},"","",""]
[null,"Bijstelling voor drift noodzakelijk","","","","",""]
[null,"Probleem met precisie; onderhoud noodzakelijk","","Geen gegevens","Geen gegevens","Geen gegevens",""]
[null,"S-AMS *","",0,0,{"$error":"#VALUE!"},"","",""]
[null,"Bijstelling voor drift noodzakelijk","","","","",""]
[null,"Probleem met precisie; onderhoud noodzakelijk","","Geen gegevens","Geen gegevens","Geen gegevens",""]
[null,"Drift correcties altijd uitgevoerd","Ja","Ja","Ja","Ja","Ja"]
[null,"Benodigd onderhoud altijd uitgevoerd","Ja","Ja","Ja","Ja","Ja"]
[]
[null," *: als standaard deviatie"]

{"sheet":"CUSUM NOx","index":15,"n_rows":178,"n_cols":29,"rows_kept":178,"cut":false,"hidden":true,"merged":[[15,2,15,14],[15,15,15,27],[15,1,16,1],[1,1,1,27],[2,1,2,27],[4,15,4,27],[4,1,4,14]],"flags":["formulas_cached_only"]}
[]
[null,"Cusum voor NOx-analyser"]
[null,"EN14181-NOx.xls Versie 2.1"]
[]
[null,"Data QAL3 procedure",{"$nulls":13},"Cusum kaart"]
[null,"Concentratie eenheid QAL3-medium"]
[null,"Correctiefactor",{"$nulls":10},2.054]
[]
[{"$nulls":12},"Nul","Span"]
[null,"S-AMS *",{"$nulls":10},"","","ppm"]
[null,"hx = 2,85 x S-AMS *",{"$nulls":10},{"$error":"#VALUE!"},{"$error":"#VALUE!"}]
[null,"kx = 0,501 x S-AMS *",{"$nulls":10},{"$error":"#VALUE!"},{"$error":"#VALUE!"}]
[null,"hs = 6,90 x (S-AMS *)^2",{"$nulls":10},{"$error":"#VALUE!"},{"$error":"#VALUE!"}]
[null,"ks = 1,85 x (S-AMS *)^2",{"$nulls":10},{"$error":"#VALUE!"},{"$error":"#VALUE!"}]
[]
[null,"Datum","Nul",{"$nulls":12},"Span"]
["Reset zero",null,"Uitlezing","Ref","d","s","∑(pos)","N(pos)","∑(neg)","N(neg)","N(s)","Berekende bijstelling voor drift","Bijstelling","Probleem met precisie; onderhoud noodzakelijk","Onderhoud uitgevoerd?","Uitlezing","Ref","d","s","∑(pos)","N(pos)","∑(neg)","N(neg)","N(s)","Berekende bijstelling voor drift","Bijstelling","Probleem met precisie; onderhoud noodzakelijk","Onderhoud uitgevoerd?","Reset span"]
["","2011-01-01",null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]

{"sheet":"CUSUM NO","index":16,"n_rows":178,"n_cols":29,"rows_kept":178,"cut":false,"hidden":true,"merged":[[15,2,15,14],[15,15,15,27],[15,1,16,1],[1,1,1,27],[2,1,2,27],[4,15,4,27],[4,1,4,14]],"flags":["formulas_cached_only"]}
[]
[null,"Cusum voor NO-analyser"]
[null,"EN14181-NOx.xls Versie 2.1"]
[]
[null,"Data QAL3 procedure",{"$nulls":13},"Cusum kaart"]
[null,"Concentratie eenheid QAL3-medium"]
[null,"Correctiefactor",{"$nulls":10},2.054]
[]
[{"$nulls":12},"Nul","Span"]
[null,"S-AMS *",{"$nulls":10},0,0,"ppm"]
[null,"hx = 2,85 x S-AMS *",{"$nulls":10},0,0]
[null,"kx = 0,501 x S-AMS *",{"$nulls":10},0,0]
[null,"hs = 6,90 x (S-AMS *)^2",{"$nulls":10},0,0]
[null,"ks = 1,85 x (S-AMS *)^2",{"$nulls":10},0,0]
[]
[null,"Datum","Nul",{"$nulls":12},"Span"]
["Reset zero",null,"Uitlezing","Ref","d","s","∑(pos)","N(pos)","∑(neg)","N(neg)","N(s)","Berekende bijstelling voor drift","Bijstelling","Probleem met precisie; onderhoud noodzakelijk","Onderhoud uitgevoerd?","Uitlezing","Ref","d","s","∑(pos)","N(pos)","∑(neg)","N(neg)","N(s)","Berekende bijstelling voor drift","Bijstelling","Probleem met precisie; onderhoud noodzakelijk","Onderhoud uitgevoerd?","Reset span"]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]

{"sheet":"CUSUM NO2","index":17,"n_rows":178,"n_cols":29,"rows_kept":178,"cut":false,"hidden":true,"merged":[[15,2,15,14],[15,15,15,27],[15,1,16,1],[1,1,1,27],[2,1,2,27],[4,15,4,27],[4,1,4,14]],"flags":["formulas_cached_only"]}
[]
[null,"Cusum voor NO2-analyser"]
[null,"EN14181-NOx.xls Versie 2.1"]
[]
[null,"Data QAL3 procedure",{"$nulls":13},"Cusum kaart"]
[null,"Concentratie eenheid QAL3-medium"]
[null,"Correctiefactor",{"$nulls":10},2.054]
[]
[{"$nulls":12},"Nul","Span"]
[null,"S-AMS *",{"$nulls":10},0,0,"ppm"]
[null,"hx = 2,85 x S-AMS *",{"$nulls":10},0,0]
[null,"kx = 0,501 x S-AMS *",{"$nulls":10},0,0]
[null,"hs = 6,90 x (S-AMS *)^2",{"$nulls":10},0,0]
[null,"ks = 1,85 x (S-AMS *)^2",{"$nulls":10},0,0]
[]
[null,"Datum","Nul",{"$nulls":12},"Span"]
["Reset zero",null,"Uitlezing","Ref","d","s","∑(pos)","N(pos)","∑(neg)","N(neg)","N(s)","Berekende bijstelling voor drift","Bijstelling","Probleem met precisie; onderhoud noodzakelijk","Onderhoud uitgevoerd?","Uitlezing","Ref","d","s","∑(pos)","N(pos)","∑(neg)","N(neg)","N(s)","Berekende bijstelling voor drift","Bijstelling","Probleem met precisie; onderhoud noodzakelijk","Onderhoud uitgevoerd?","Reset span"]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]

{"sheet":"CUSUM O2","index":18,"n_rows":178,"n_cols":29,"rows_kept":178,"cut":false,"hidden":true,"merged":[[15,2,15,14],[15,15,15,27],[15,1,16,1],[1,1,1,27],[2,1,2,27],[4,15,4,27],[4,1,4,14]],"flags":["formulas_cached_only"]}
[]
[null,"Cusum voor O2-analyser"]
[null,"EN14181-NOx.xls Versie 2.1"]
[]
[null,"Data QAL3 procedure",{"$nulls":13},"Cusum kaart"]
[null,"Concentratie eenheid QAL3-medium"]
[null,"Correctiefactor",{"$nulls":10},1]
[]
[{"$nulls":12},"Nul","Span"]
[null,"S-AMS *",{"$nulls":10},{"$error":"#VALUE!"},{"$error":"#VALUE!"},"vol%"]
[null,"hx = 2,85 x S-AMS *",{"$nulls":10},{"$error":"#VALUE!"},{"$error":"#VALUE!"}]
[null,"kx = 0,501 x S-AMS *",{"$nulls":10},{"$error":"#VALUE!"},{"$error":"#VALUE!"}]
[null,"hs = 6,90 x (S-AMS *)^2",{"$nulls":10},{"$error":"#VALUE!"},{"$error":"#VALUE!"}]
[null,"ks = 1,85 x (S-AMS *)^2",{"$nulls":10},{"$error":"#VALUE!"},{"$error":"#VALUE!"}]
[]
[null,"Datum","Nul",{"$nulls":12},"Span"]
["Reset zero",null,"Uitlezing","Ref","d","s","∑(pos)","N(pos)","∑(neg)","N(neg)","N(s)","Berekende bijstelling voor drift","Bijstelling","Probleem met precisie; onderhoud noodzakelijk","Onderhoud uitgevoerd?","Uitlezing","Ref","d","s","∑(pos)","N(pos)","∑(neg)","N(neg)","N(s)","Berekende bijstelling voor drift","Bijstelling","Probleem met precisie; onderhoud noodzakelijk","Onderhoud uitgevoerd?","Reset span"]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]

{"sheet":"CUSUM H2O","index":19,"n_rows":178,"n_cols":29,"rows_kept":178,"cut":false,"hidden":true,"merged":[[15,2,15,14],[15,15,15,27],[15,1,16,1],[1,1,1,27],[2,1,2,27],[4,15,4,27],[4,1,4,14]],"flags":["formulas_cached_only"]}
[]
[null,"Cusum voor H2O-analyser"]
[null,"EN14181-NOx.xls Versie 2.1"]
[]
[null,"Data QAL3 procedure",{"$nulls":13},"Cusum kaart"]
[null,"Concentratie eenheid QAL3-medium"]
[null,"Correctiefactor",{"$nulls":10},1]
[]
[{"$nulls":12},"Nul","Span"]
[null,"S-AMS *",{"$nulls":10},"","","vol%"]
[null,"hx = 2,85 x S-AMS *",{"$nulls":10},{"$error":"#VALUE!"},{"$error":"#VALUE!"}]
[null,"kx = 0,501 x S-AMS *",{"$nulls":10},{"$error":"#VALUE!"},{"$error":"#VALUE!"}]
[null,"hs = 6,90 x (S-AMS *)^2",{"$nulls":10},{"$error":"#VALUE!"},{"$error":"#VALUE!"}]
[null,"ks = 1,85 x (S-AMS *)^2",{"$nulls":10},{"$error":"#VALUE!"},{"$error":"#VALUE!"}]
[]
[null,"Datum","Nul",{"$nulls":12},"Span"]
["Reset zero",null,"Uitlezing","Ref","d","s","∑(pos)","N(pos)","∑(neg)","N(neg)","N(s)","Berekende bijstelling voor drift","Bijstelling","Probleem met precisie; onderhoud noodzakelijk","Onderhoud uitgevoerd?","Uitlezing","Ref","d","s","∑(pos)","N(pos)","∑(neg)","N(neg)","N(s)","Berekende bijstelling voor drift","Bijstelling","Probleem met precisie; onderhoud noodzakelijk","Onderhoud uitgevoerd?","Reset span"]
["","2011-01-01",null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,null,"","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,null,"","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]
["",null,null,"","","","","","","","","",null,"",null,null,"","","","","","","","","",null,"",null,""]

{"sheet":"AST","index":20,"n_rows":15,"n_cols":5,"rows_kept":15,"cut":false,"hidden":true,"merged":[[1,1,1,6],[2,1,2,6]],"flags":["formulas_cached_only"]}
[]
[null,"Resultaten AST voor NOx-meetsystemen"]
[null,"EN14181-NOx.xls Versie 2.1"]
[]
[null,"Parallelmetingen (Jaar 2)"]
[null,"Datum","","",""]
[null,"Meetinstantie",""]
[null,"Aantal parallelmetingen",0]
[null,"Variabiliteitstest",{"$error":"#NUM!"}]
[null,"Test voor validiteit kalibratie lijn","Niet-valide"]
[null,"Opmerkingen","Te weinig metingen!"]
[]
[null,"Functionele testen (Jaar 2)"]
[null,"Datum","","",""]
[null,"Tests uitgevoerd","Niet allemaal"]

{"sheet":"AST parallel measurements 1","index":21,"n_rows":58,"n_cols":17,"rows_kept":58,"cut":false,"hidden":true,"merged":[[1,1,1,10],[2,1,2,10],[5,6,5,10],[6,6,6,10],[16,6,16,10],[17,6,17,10],[23,4,23,10],[26,2,26,3],[27,2,27,3],[23,1,24,2],[25,2,25,3],[18,6,18,10],[19,6,19,10],[32,2,32,3],[33,2,33,3],[34,2,34,3],[24,4,24,6],[28,2,28,3],[24,7,24,9],[42,2,42,3],[24,13,25,13],[24,16,25,16],[24,12,25,12],[24,14,25,14],[24,15,25,15],[38,2,38,3],[29,2,29,3],[30,2,30,3],[31,2,31,3],[39,2,39,3],[40,2,40,3],[41,2,41,3],[35,2,35,3],[36,2,36,3],[37,2,37,3]],"flags":["formulas_cached_only"]}
[]
[null,"AST-procedure van het NOx-meetsysteem "]
[null,"EN14181-NOx.xls Versie 2.1"]
[]
[null,"Specificatie meetsysteem"]
[null,"NOx-meetmethode",{"$nulls":4},"Extractieve NO- + NO2-meting in droog rookgas"]
[null,"O2-meetmethode",{"$nulls":4},"In-situ O2-meting"]
[null,"Kalibratielijn gebruikt tijdens AST",{"$nulls":4},"QAL2","2015-01-01"]
[null,null,"Intercept",null,null,null,""]
[null,null,"Helling",null,null,null,""]
[null,null,"Geldig kalibratiebereik",null,null,null,"","mg/Nm3 (als NO2) droog rookgas bij  vol% O2"]
[null,"Emissiegrenswaarde",{"$nulls":4},0,"mg/Nm3 (als NO2) droog rookgas bij  vol% O2"]
[null,"Testwaarde bij standaard condities",{"$nulls":4},0,"mg/Nm3 (als NO2) droog rookgas bij  vol% O2"]
[null,"Vereist betrouwbaarheidsinterval",{"$nulls":4},0,"%"]
[]
[null,"Standaard referentie methode (SRM)"]
[null,"Meetinstantie"]
[null,"NOx-meetmethode"]
[null,"O2-meetmethode"]
[null,"Meetplaats"]
[null,"Monsternamemethode"]
[null,"Aantal meetpunten SRM"]
[]
[null,"AST meetpunten",null,null,"NOx-concentratie bij standaard condities (mg/Nm3 (als NO2) droog rookgas bij  vol% O2)"]
[{"$nulls":4},"Standaard referentie methode (SRM)",null,null,"NOx-meetsysteem (AMS)",null,null,"Δ",null,"AMS","SRM","Δ","(Δ-Δgem)2","Dag"]
[null,"Nummer","Datum",null,"Starttijd","Eindtijd","Meetwaarde","Starttijd","Eindtijd","Meetwaarde","AMS - SRM"]
[null,"",{"$nulls":8},"",null,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"Min",{"$nulls":4},"",null,null,"",{"$nulls":4},"",""]
[null,"Max","",null,null,null,"",null,null,"",{"$nulls":4},"","",0]
[null,"Gemiddeld",{"$nulls":4},"",null,null,"",{"$nulls":4},"",""]
[]
[null,"AST resultaten"]
[null,"Aantal parallelmetingen",null,null,null,0,"Te weinig metingen!"]
[null,"Variabiliteitstest"]
[null,null,"S0",null,null,0]
[null,null,"KV",null,null,{"$error":"#NUM!"}]
[null,null,"1,5 x KV x S0 ",null,null,{"$error":"#NUM!"}]
[null,null,"SD",null,null,""]
[null,null,"1,5 x KV x S0 ",{"$error":"#NUM!"},"SD",{"$error":"#NUM!"}]
[null,"Test voor validiteit kalibratie lijn"]
[null,null,"t0,95",null,null,{"$error":"#NUM!"}]
[null,null,"│D│",null,null,""]
[null,null,"t0,95 x SD / √N + S0",null,null,""]
[null,null,"t0,95 x SD / √N + S0","<=","│D│","Niet-valide"]

{"sheet":"AST functionality 1","index":22,"n_rows":16,"n_cols":18,"rows_kept":16,"cut":false,"hidden":true,"merged":[[1,1,1,5],[2,1,2,5]],"flags":["formulas_cached_only"]}
[]
[null,"AST - functionele testen"]
[null,"EN14181-NOx.xls Versie 2.1"]
[]
[null,"Test","Datum","Uitvoerder: naam / bedrijf","Methode","Resultaat",0,"test","dat","uitv","meth","res","som","ok1","ok2",null,"tot"]
[null,"Uitlijning en vervuiling",{"$nulls":4},1,1,0,0,0,0,0,0,0,null,0,0]
[null,"Monsternamesysteem",{"$nulls":4},2,1,0,0,0,0,0,0,0,null,0]
[null,"Documentatie en verslagen",{"$nulls":4},3,1,0,0,0,0,0,0,0,null,0]
[null,"Onderhoudgeschiktheid",{"$nulls":4},4,1,0,0,0,0,0,0,0,null,0]
[null,"Controle van nul en span",{"$nulls":4},5,1,0,0,0,0,0,0,0,null,0]
[null,"Lektoets",{"$nulls":4},6,1,0,0,0,0,0,0,0,null,0]
[null,"Lineariteit",{"$nulls":4},7,1,0,0,0,0,0,0,0,null,0]
[null,"Invloed storende factoren",{"$nulls":4},8,1,0,0,0,0,0,0,0,null,0]
[null,"Drift van nul en span (audit)",{"$nulls":4},9,1,0,0,0,0,0,0,0,null,0]
[null,"Responstijd",{"$nulls":4},10,1,0,0,0,0,0,0,0,null,0]
[null,"Rapportage",{"$nulls":4},11,1,0,0,0,0,0,0,0,null,0]

{"sheet":"AST parallel measurements 2","index":23,"n_rows":58,"n_cols":17,"rows_kept":58,"cut":false,"hidden":true,"merged":[[24,16,25,16],[24,12,25,12],[24,14,25,14],[24,15,25,15],[34,2,34,3],[39,2,39,3],[35,2,35,3],[36,2,36,3],[37,2,37,3],[25,2,25,3],[24,13,25,13],[38,2,38,3],[29,2,29,3],[30,2,30,3],[31,2,31,3],[32,2,32,3],[33,2,33,3],[24,4,24,6],[19,6,19,10],[23,4,23,10],[26,2,26,3],[24,7,24,9],[27,2,27,3],[42,2,42,3],[40,2,40,3],[41,2,41,3],[28,2,28,3],[23,1,24,2],[18,6,18,10],[1,1,1,10],[2,1,2,10],[5,6,5,10],[6,6,6,10],[16,6,16,10],[17,6,17,10]],"flags":["formulas_cached_only"]}
[]
[null,"AST-procedure van het NOx-meetsysteem "]
[null,"EN14181-NOx.xls Versie 2.1"]
[]
[null,"Specificatie meetsysteem"]
[null,"NOx-meetmethode",{"$nulls":4},"Extractieve NO- + NO2-meting in droog rookgas"]
[null,"O2-meetmethode",{"$nulls":4},"In-situ O2-meting"]
[null,"Kalibratielijn gebruikt tijdens AST",{"$nulls":4},"QAL2","2015-01-01"]
[null,null,"Intercept",null,null,null,""]
[null,null,"Helling",null,null,null,""]
[null,null,"Geldig kalibratiebereik",null,null,null,"","mg/Nm3 (als NO2) droog rookgas bij  vol% O2"]
[null,"Emissiegrenswaarde",{"$nulls":4},0,"mg/Nm3 (als NO2) droog rookgas bij  vol% O2"]
[null,"Testwaarde bij standaard condities",{"$nulls":4},0,"mg/Nm3 (als NO2) droog rookgas bij  vol% O2"]
[null,"Vereist betrouwbaarheidsinterval",{"$nulls":4},0,"%"]
[]
[null,"Standaard referentie methode (SRM)"]
[null,"Meetinstantie"]
[null,"NOx-meetmethode"]
[null,"O2-meetmethode"]
[null,"Meetplaats"]
[null,"Monsternamemethode"]
[null,"Aantal meetpunten SRM"]
[]
[null,"AST meetpunten",null,null,"NOx-concentratie bij standaard condities (mg/Nm3 (als NO2) droog rookgas bij  vol% O2)"]
[{"$nulls":4},"Standaard referentie methode (SRM)",null,null,"NOx-meetsysteem (AMS)",null,null,"Δ",null,"AMS","SRM","Δ","(Δ-Δgem)2","Dag"]
[null,"Nummer","Datum",null,"Starttijd","Eindtijd","Meetwaarde","Starttijd","Eindtijd","Meetwaarde","AMS - SRM"]
[null,"",{"$nulls":8},"",null,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"",{"$nulls":8},"",0,"","","","",0]
[null,"Min",{"$nulls":4},"",null,null,"",{"$nulls":4},"",""]
[null,"Max","",null,null,null,"",null,null,"",{"$nulls":4},"","",0]
[null,"Gemiddeld",{"$nulls":4},"",null,null,"",{"$nulls":4},"",""]
[]
[null,"AST resultaten"]
[null,"Aantal parallelmetingen",null,null,null,0,"Te weinig metingen!"]
[null,"Variabiliteitstest"]
[null,null,"S0",null,null,0]
[null,null,"KV",null,null,{"$error":"#NUM!"}]
[null,null,"1,5 x KV x S0 ",null,null,{"$error":"#NUM!"}]
[null,null,"SD",null,null,""]
[null,null,"1,5 x KV x S0 ",{"$error":"#NUM!"},"SD",{"$error":"#NUM!"}]
[null,"Test voor validiteit kalibratie lijn"]
[null,null,"t0,95",null,null,{"$error":"#NUM!"}]
[null,null,"│D│",null,null,""]
[null,null,"t0,95 x SD / √N + S0",null,null,""]
[null,null,"t0,95 x SD / √N + S0","<=","│D│","Niet-valide"]

{"sheet":"AST functionality 2","index":24,"n_rows":16,"n_cols":18,"rows_kept":16,"cut":false,"hidden":true,"merged":[[1,1,1,5],[2,1,2,5]],"flags":["formulas_cached_only"]}
[]
[null,"AST - functionele testen"]
[null,"EN14181-NOx.xls Versie 2.1"]
[]
[null,"Test","Datum","Uitvoerder: naam / bedrijf","Methode","Resultaat",0,"test","dat","uitv","meth","res","som","ok1","ok2",null,"tot"]
[null,"Uitlijning en vervuiling",{"$nulls":4},1,1,0,0,0,0,0,0,0,null,0,0]
[null,"Monsternamesysteem",{"$nulls":4},2,1,0,0,0,0,0,0,0,null,0]
[null,"Documentatie en verslagen",{"$nulls":4},3,1,0,0,0,0,0,0,0,null,0]
[null,"Onderhoudgeschiktheid",{"$nulls":4},4,1,0,0,0,0,0,0,0,null,0]
[null,"Controle van nul en span",{"$nulls":4},5,1,0,0,0,0,0,0,0,null,0]
[null,"Lektoets",{"$nulls":4},6,1,0,0,0,0,0,0,0,null,0]
[null,"Lineariteit",{"$nulls":4},7,1,0,0,0,0,0,0,0,null,0]
[null,"Invloed storende factoren",{"$nulls":4},8,1,0,0,0,0,0,0,0,null,0]
[null,"Drift van nul en span (audit)",{"$nulls":4},9,1,0,0,0,0,0,0,0,null,0]
[null,"Responstijd",{"$nulls":4},10,1,0,0,0,0,0,0,0,null,0]
[null,"Rapportage",{"$nulls":4},11,1,0,0,0,0,0,0,0,null,0]

{"sheet":"Blad1","index":25,"n_rows":0,"n_cols":0,"rows_kept":0,"cut":false,"hidden":true,"merged":[],"flags":["formulas_cached_only"]}